## Version History
Version | Changes | Author | Reviewer | Date | External Firmware Version
v1.01 | Initial release | 王少英 |  | 2024-07-26
v1.1.2 | (1) Add boot loader in System Information (2) RS485 parity default value is changed to 3(none1) | 王少英 |  | 2024-11-18
v1.13 | 1. In the sheet"Overview" add the address description of "Real-Time Parameters (float)-High Speed" 2. In the sheet"Overview" add the address description of "Real-Time Parameters (int)-High Speed" 3. In the sheet "Readings",add the register map of "Real-time Parameters (float) -High Speed" 4. In the sheet "Readings" add the register map of "Real-Time Parameters (int)-High Speed" 5. In the sheet "System Information",add the register "Bootloader Patch Number" 6. In the sheet "System Information",add the register "Hardwaer Patch Number" | 王培功 |  | 2025-05-13
v1.1.4 | 1.In the sheet"System Information" : Change the descritpion of [F040 - F04F] 2.In the sheet"System Information": Change the register address of Hardware Patch number | 王培功 |  | 2025-05-30
v1.1.5 | 1. In the sheet "Readings", add the register of temperature and the reserved | 王培功 |  | 2025-06-20
v1.1.6 | (1)"TransactionLog" Sheet is added (2) Add "Clear Transaction Log" in Clear & Reset (3) Add "public key", "private key", "factory test sign", "factory test sign size in bytes", "CRC32 for product info"，“System Mode”， “System Status" to [Product Info] sections in [System Infomation] Sheet | 王少英 |  | 2025-08-14
v1.1.7 | 1.In the sheet"Basic Configuration": Add the "Measurement data integrity check CRC32" register. 2.In the sheet"Basic Configuration": Add the "Clear Echilog" register. 3.In the sheet"Readings": Add the "Energy Data Integrity Check CRC32" register. 4.In the sheet"System Information": Add the data block "Meter System Status" 5.Add the sheet "Echilog" 6.In the sheet "Calibration",add the "Calibration Integrity Check CRC32" register,and change the ADC calibration register "WR" to "RO" | 王培功 |  | 2025-08-21
v1.1.8 | 1. Add [Meter System Status] in [Overview] sheet 2. Update [Function Model Type] in [Product info] section, [System Information] sheet | 王少英 |  | 2025-08-22
v1.1.9 | (1) [default] and [factory value] of [Function Model Type] in [System Information] sheet is set to 1 for DC300 Sprint2 | Wang Shaoying |  | 2025-09-11
v1.2.0 | (1) Add [Remote Display Mode] to [Product Info] in [System Information] sheet | Wang Shaoying |  | 2025-10-31
v1.2.1 | (1)Change length of public key from 65 bytes to 91 bytes | Wang Shaoying |  | 2025-11-12
v1.2.2 | (1)Delete[DHCP] , IP, Subnetmask, gateway, DNS1, DNS2 in [Basic Configuration] sheet (2) Delete[Modbus Tcp Enable] and [Modbus Tcp Port] in [Basic Configuration] Sheet (3) Delete [Network Reset] in [Clear & Reset] in [Basic Configuration] sheet | Wang Shaoying |  | 2025-11-13
v1.2.3 | Change values of PT1,PT2,CT1,CT2 in [Basic Meter Settings] seciton of [Basic Configuration] sheet | Wang Shaoying |  | 2025-12-05

## Overview
Block | Sheet | Start | End | Length | Comments
Basic Meter Settings | Basic Configuration | 0x1000 | 0x1023 | 36
Real-Time Clock | Basic Configuration | 0x103F | 0x1045 | 7
Datalog 1, 2, 3 Settings | Datalog | 0x1100 | 0x1178 | 121
Datalog 1 Sector | Datalog | 0x1400 | 0x1402 | 3
Datalog 2 Sector | Datalog | 0x1410 | 0x1412 | 3
Datalog 3 Sector | Datalog | 0x1420 | 0x1422 | 3
Datalog 4 Sector | Datalog | 0x1430 | 0x1432 | 3
Datalog 4 Settings | Datalog | 0x1500 | 0x15CC | 205
Clear & Reset | Basic Configuration | 0x2000 | 0x2009 | 10
Real-Time Parameters (float) | Readings | 0x3000 | 0x3021 | 34
Real-Time Parameters (float)-High Speed | Readings | 0x3100 | 0x3109 | 10
Real-Time Parameters (int) | Readings | 0x3200 | 0x3210 | 17
Real-Time Parameters (int)-High Speed | Readings | 0x3300 | 0x3304 | 5
Energy and Charge | Readings | 0x4000 | 0x401F | 32
Max | Max/Min | 0x4136 | 0x4162 | 45
Demand Max | Max/Min | 0x4163 | 0x4186 | 36
Min | Max/Min | 0x4187 | 0x41B3 | 45
Setting | TransactionLog | 0x5000 | 0x5049 | 74
Block and Id | TransactionLog | 0x5100 | 0x5106 | 7
Reading | TransactionLog | 0x5200 | 0x5478 | 633
Status | TransactionLog | 0x5500 | 0x5504 | 5
Status | Echilog | 0x5700 | 0x5702 | 3
Reading | Echilog | 0x5800 | 0x587B | 124
Datalog 1, 2, 3 Reading | Datalog | 0x6000 | 0x60CC | 205
Datalog 1 Status | Datalog | 0x6100 | 0x610D | 14
Datalog 2 Status | Datalog | 0x6200 | 0x620D | 14
Datalog 3 Status | Datalog | 0x6300 | 0x630D | 14
Datalog 4 Reading | Datalog | 0x6400 | 0x64CC | 205
Datalog 4 Status | Datalog | 0x6500 | 0x650D | 14
Custom Reading | Basic Configuration | 0x6A00 | 0x6AC7 | 200
Custom Reading Settings | Basic Configuration | 0x6B00 | 0x6B32 | 51
Firmware Version Info | System Information | 0xF000 | 0xF005 | 6
Boot Loader | System Information | 0xF010 | 0xF017 | 8
Product Info | System Information | 0xF040 | 0xF0AD | 110
Meter System Status | System Information | 0xF0B4 | 0xF0B7 | 4
Public key | System Information | 0xF0C0 | 0xF0ED | 46
Calibration Commands | Calibration | 0xF204 | 0xF209 | 6
Calibration Values | Calibration | 0xF210 | 0xF217 | 8
废弃
废弃
废弃

## Basic Configuration
Block | Start(Hex) | End(Hex) | Start(Dec) | End(Dec) | Descrption | Data type | RW | Reg | RANGE | Default | Factory value
Basic Meter Settings | 0x1000 | 0x1000 | 4096 | 4096 | Meter password | word | R/W | 1 | 0－9999 | 0
| 0x1001 | 0x1001 | 4097 | 4097 | RS485 baud rate | word | R/W | 1 | 2400－115200 | 19200
| 0x1002 | 0x1002 | 4098 | 4098 | RS485 parity | word | R/W | 1 | 0：even 1：odd 2：Non2 3：Non1 | 3
| 0x1003 | 0x1003 | 4099 | 4099 | DHCP enable | word | R/W | 1 | 0: manual 1: DHCP | 0
| 0x1004 | 0x1004 | 4100 | 4100 | IP address 1st byte (high) IP address 2nd byte (low) | word | R/W | 1 | 0~255 |  | 192 168 1 254
| 0x1005 | 0x1005 | 4101 | 4101 | IP address 3rd byte (high) IP address 4th byte (low) | word | R/W | 1 | 0~255
| 0x1006 | 0x1006 | 4102 | 4102 | Subnet mask 1st byte (high) Subnet mask 2nd byte (low) | word | R/W | 1 | 0~255 |  | 255 255 255 0
| 0x1007 | 0x1007 | 4103 | 4103 | Subnet mask 3rd byte (high) Subnet mask 4th byte (low) | word | R/W | 1 | 0~255
| 0x1008 | 0x1008 | 4104 | 4104 | Gateway 1st byte (high) Gateway 2nd byte (low) | word | R/W | 1 | 0~255 |  | 192 168 1 1
| 0x1009 | 0x1009 | 4105 | 4105 | Gateway 3rd byte (high) Gateway 4th byte (low) | word | R/W | 1 | 0~255
| 0x100A | 0x100A | 4106 | 4106 | DNS primary server 1st byte (high) DNS primary server 2nd byte (low) | word | R/W | 1 | 0~255 |  | 8 8 8 8
| 0x100B | 0x100B | 4107 | 4107 | DNS primary server 3rd byte (high) DNS primary server 4th byte (low) | word | R/W | 1 | 0~255
| 0x100C | 0x100C | 4108 | 4108 | DNS secondary server 1st byte (high) DNS secondary server 2nd byte (low) | word | R/W | 1 | 0~255 |  | 8.8.4.4
| 0x100D | 0x100D | 4109 | 4109 | DNS secondary server 3rd byte (high) DNS secondary server 4th byte (low) | word | R/W | 1 | 0~255
| 0x100E | 0x100E | 4110 | 4110 | Modbus Slave ID | word | R/W | 1 | 1-247, used for both Modbus RTU and TCP | 1 | 1
| 0x100F | 0x100F | 4111 | 4111 | Modbus RTU Enable | word | R/W | 1 | 0: disable 1: enable | 1 | 1
| 0x1010 | 0x1010 | 4112 | 4112 | Modbus TCP Enable | word | R/W | 1 | 0: disable 1: enable | 1 | 1
| 0x1011 | 0x1011 | 4113 | 4113 | Modbus TCP port | word | R/W | 1 | 1-65534 | 502 | 502
| 0x1012 | 0x1012 | 4114 | 4114 | PT1 | word | R | 1 | fixed as 1500 for now | 1500
| 0x1013 | 0x1013 | 4115 | 4115 | PT2 | word | R | 1 | fixed as 1500 for now | 1500
| 0x1014 | 0x1014 | 4116 | 4116 | CT1 | word | R | 1 | fixed as 1500A | 1500
| 0x1015 | 0x1015 | 4117 | 4117 | CT2 | word | R | 1 | fixed as 15mV | 15
| 0x1016 | 0x1016 | 4118 | 4118 | Demand calculation method | word | R/W | 1 | 0: Fixed block 1: Sliding block | 1
| 0x1017 | 0x1017 | 4119 | 4119 | Demand window time | word | R/W | 1 | 1~30 | 15
| 0x1018 | 0x1018 | 4120 | 4120 | Demand update period | word | R/W | 1 | 1~30 | 1
| 0x1019 | 0x1019 | 4121 | 4121 | Energy pulse parameter | word | R/W | 1 | 0: None 1: Import Energy 2: Export Energy 3: NET Energy 4: TOTAL Energy | 0
| 0x101A | 0x101B | 4122 | 4123 | Energy pulse constant | word | R/W | 2 | 0.1—100000.000 | 1 |  | internal actual value is *1000 in integers.
| 0x101C | 0x101C | 4124 | 4124 | Backlight time | word | R/W | 1 | 0－120 minutes | 1
| 0x101D | 0x101D | 4125 | 4125 | Seal status | word | R | 1 | Sealed: 0x0A | 0
| 0x101E | 0x101E | 4126 | 4126 | Device run time (high 16 bits) | word | R/W | 1 | uint in second
| 0x101F | 0x101F | 4127 | 4127 | Device run time (low 16 bits) | Word | R/W | 1
| 0x1020 | 0x1020 | 4128 | 4128 | Device load time (high 16 bits) | word | R/W | 1 | uint in second
| 0x1021 | 0x1021 | 4129 | 4129 | Device load time (low 16 bits) | word | R/W | 1
| 0x1022 | 0x1022 | 4130 | 4130 | Enable cable loss compensation | word | R/W | 1 | 0: Disable 1: Enable | 0
| 0x1023 | 0x1023 | 4131 | 4131 | Cable resistance | word | R/W | 1 | 0 ~ 65535(in 0.0001 Ohm) | 0
| 0x1024 | 0x1025 | 4132 | 4133 | Measurement data integrity check CRC32 | uint16_t | RO | 2 | * Not available for Up Computer | 0
Real-time clock
| 0x103F | 0x103F | 4159 | 4159 | Week | word | R/W | 1 | 1~7 | 1
| 0x1040 | 0x1040 | 4160 | 4160 | Year | word | R/W | 1 | 2000~2099 | 2024
| 0x1041 | 0x1041 | 4161 | 4161 | Month | word | R/W | 1 | 1~12 | 1
| 0x1042 | 0x1042 | 4162 | 4162 | Day | word | R/W | 1 | 1~31 | 1
| 0x1043 | 0x1043 | 4163 | 4163 | Hour | word | R/W | 1 | 0~23 | 0
| 0x1044 | 0x1044 | 4164 | 4164 | Minute | word | R/W | 1 | 0~59 | 0
| 0x1045 | 0x1045 | 4165 | 4165 | Second | word | R/W | 1 | 0~59 | 0
Clear & Reset | 0x2000 | 0x2000 | 8192 | 8192 | Clear energy | word | R/W | 1 | 0: None 1: Clearing | 0
| 0x2001 | 0x2001 | 8193 | 8193 | Clear charge | word | R/W | 1 | 0: None 1: Clearing | 0
| 0x2002 | 0x2002 | 8194 | 8194 | Clear demand | word | R/W | 1 | 0: None 1: Clearing | 0
| 0x2003 | 0x2003 | 8195 | 8195 | Clear max/min | word | R/W | 1 | 0: None 1: Clearing | 0
| 0x2004 | 0x2004 | 8196 | 8196 | Clear device run time | word | R/W | 1 | 0: None 1: Clearing | 0
| 0x2005 | 0x2005 | 8197 | 8197 | Clear load time | word | R/W | 1 | 0: None 1: Clearing | 0
| 0x2006 | 0x2006 | 8198 | 8198 | Factory reset | word | R/W | 1 | 0: None 1: Resetting | 0
| 0x2007 | 0x2007 | 8199 | 8199 | Network reset | word | R/W | 1 | 0: None 1: Resetting | 0
| 0x2008 | 0x2008 | 8200 | 8200 | Clear Transaction Log | uint16 | RW | 1 | 0: None 1: Clearing
| 0x2009 | 0x2009 | 8201 | 8201 | Clear Echilog | uint16_t | RW | 1 | 0: None 1: Clearing
Custom reading
| 0x6A00 | 0x6A00 | 27136 | 27136 | 1st reading |  | R | 1 |  |  | 0
| 0x6A01 | 0x6A01 | 27137 | 27137 | 2nd reading |  | R | 1 |  |  | 0
| 0x6A02 | 0x6A02 | 27138 | 27138 | 3rd reading |  | R | 1 |  |  | 0
| 0x6A03 | 0x6A03 | 27139 | 27139 | 4th reading |  | R | 1 |  |  | 0
| 0x6A04 | 0x6A04 | 27140 | 27140 | 5th reading |  | R | 1 |  |  | 0
| 0x6A05 | 0x6A05 | 27141 | 27141 | 6th reading |  | R | 1 |  |  | 0
| 0x6A06 | 0x6A06 | 27142 | 27142 | 7th reading |  | R | 1 |  |  | 0
| 0x6A07 | 0x6A07 | 27143 | 27143 | 8th reading |  | R | 1 |  |  | 0
| 0x6A08 | 0x6A08 | 27144 | 27144 | 9th reading |  | R | 1 |  |  | 0
| 0x6A09 | 0x6A09 | 27145 | 27145 | 10th reading |  | R | 1 |  |  | 0
| 0x6A0A | 0x6A0A | 27146 | 27146 | 11th reading |  | R | 1 |  |  | 0
| 0x6A0B | 0x6A0B | 27147 | 27147 | 12th reading |  | R | 1 |  |  | 0
| 0x6A0C | 0x6A0C | 27148 | 27148 | 13th reading |  | R | 1 |  |  | 0
| 0x6A0D | 0x6A0D | 27149 | 27149 | 14th reading |  | R | 1 |  |  | 0
| 0x6A0E | 0x6A0E | 27150 | 27150 | 15th reading |  | R | 1 |  |  | 0
| 0x6A0F | 0x6A0F | 27151 | 27151 | 16th reading |  | R | 1 |  |  | 0
| 0x6A10 | 0x6A10 | 27152 | 27152 | 17th reading |  | R | 1 |  |  | 0
| 0x6A11 | 0x6A11 | 27153 | 27153 | 18th reading |  | R | 1 |  |  | 0
| 0x6A12 | 0x6A12 | 27154 | 27154 | 19th reading |  | R | 1 |  |  | 0
| 0x6A13 | 0x6A13 | 27155 | 27155 | 20th reading |  | R | 1 |  |  | 0
| 0x6A14 | 0x6A14 | 27156 | 27156 | 21st reading |  | R | 1 |  |  | 0
| 0x6A15 | 0x6A15 | 27157 | 27157 | 22nd reading |  | R | 1 |  |  | 0
| 0x6A16 | 0x6A16 | 27158 | 27158 | 23rd reading |  | R | 1 |  |  | 0
| 0x6A17 | 0x6A17 | 27159 | 27159 | 24th reading |  | R | 1 |  |  | 0
| 0x6A18 | 0x6A18 | 27160 | 27160 | 25th reading |  | R | 1 |  |  | 0
| 0x6A19 | 0x6A19 | 27161 | 27161 | 26th reading |  | R | 1 |  |  | 0
| 0x6A1A | 0x6A1A | 27162 | 27162 | 27th reading |  | R | 1 |  |  | 0
| 0x6A1B | 0x6A1B | 27163 | 27163 | 28th reading |  | R | 1 |  |  | 0
| 0x6A1C | 0x6A1C | 27164 | 27164 | 29th reading |  | R | 1 |  |  | 0
| 0x6A1D | 0x6A1D | 27165 | 27165 | 30th reading |  | R | 1 |  |  | 0
| 0x6A1E | 0x6A1E | 27166 | 27166 | 31st reading |  | R | 1 |  |  | 0
| 0x6A1F | 0x6A1F | 27167 | 27167 | 32nd reading |  | R | 1 |  |  | 0
| 0x6A20 | 0x6A20 | 27168 | 27168 | 33rd reading |  | R | 1 |  |  | 0
| 0x6A21 | 0x6A21 | 27169 | 27169 | 34th reading |  | R | 1 |  |  | 0
| 0x6A22 | 0x6A22 | 27170 | 27170 | 35th reading |  | R | 1 |  |  | 0
| 0x6A23 | 0x6A23 | 27171 | 27171 | 36th reading |  | R | 1 |  |  | 0
| 0x6A24 | 0x6A24 | 27172 | 27172 | 37th reading |  | R | 1 |  |  | 0
| 0x6A25 | 0x6A25 | 27173 | 27173 | 38th reading |  | R | 1 |  |  | 0
| 0x6A26 | 0x6A26 | 27174 | 27174 | 39th reading |  | R | 1 |  |  | 0
| 0x6A27 | 0x6A27 | 27175 | 27175 | 40th reading |  | R | 1 |  |  | 0
| 0x6A28 | 0x6A28 | 27176 | 27176 | 41st reading |  | R | 1 |  |  | 0
| 0x6A29 | 0x6A29 | 27177 | 27177 | 42nd reading |  | R | 1 |  |  | 0
| 0x6A2A | 0x6A2A | 27178 | 27178 | 43rd reading |  | R | 1 |  |  | 0
| 0x6A2B | 0x6A2B | 27179 | 27179 | 44th reading |  | R | 1 |  |  | 0
| 0x6A2C | 0x6A2C | 27180 | 27180 | 45th reading |  | R | 1 |  |  | 0
| 0x6A2D | 0x6A2D | 27181 | 27181 | 46th reading |  | R | 1 |  |  | 0
| 0x6A2E | 0x6A2E | 27182 | 27182 | 47th reading |  | R | 1 |  |  | 0
| 0x6A2F | 0x6A2F | 27183 | 27183 | 48th reading |  | R | 1 |  |  | 0
| 0x6A30 | 0x6A30 | 27184 | 27184 | 49th reading |  | R | 1 |  |  | 0
| 0x6A31 | 0x6A31 | 27185 | 27185 | 50th reading |  | R | 1 |  |  | 0
| 0x6A32 | 0x6A32 | 27186 | 27186 | 51st reading |  | R | 1 |  |  | 0
| 0x6A33 | 0x6A33 | 27187 | 27187 | 52nd reading |  | R | 1 |  |  | 0
| 0x6A34 | 0x6A34 | 27188 | 27188 | 53rd reading |  | R | 1 |  |  | 0
| 0x6A35 | 0x6A35 | 27189 | 27189 | 54th reading |  | R | 1 |  |  | 0
| 0x6A36 | 0x6A36 | 27190 | 27190 | 55th reading |  | R | 1 |  |  | 0
| 0x6A37 | 0x6A37 | 27191 | 27191 | 56th reading |  | R | 1 |  |  | 0
| 0x6A38 | 0x6A38 | 27192 | 27192 | 57th reading |  | R | 1 |  |  | 0
| 0x6A39 | 0x6A39 | 27193 | 27193 | 58th reading |  | R | 1 |  |  | 0
| 0x6A3A | 0x6A3A | 27194 | 27194 | 59th reading |  | R | 1 |  |  | 0
| 0x6A3B | 0x6A3B | 27195 | 27195 | 60th reading |  | R | 1 |  |  | 0
| 0x6A3C | 0x6A3C | 27196 | 27196 | 61st reading |  | R | 1 |  |  | 0
| 0x6A3D | 0x6A3D | 27197 | 27197 | 62nd reading |  | R | 1 |  |  | 0
| 0x6A3E | 0x6A3E | 27198 | 27198 | 63rd reading |  | R | 1 |  |  | 0
| 0x6A3F | 0x6A3F | 27199 | 27199 | 64th reading |  | R | 1 |  |  | 0
| 0x6A40 | 0x6A40 | 27200 | 27200 | 65th reading |  | R | 1 |  |  | 0
66th-200th are reserved. | 0x6A41 | 0x6A41 | 27201 | 27201 | 66th reading |  | R | 1 |  |  | 0
| 0x6A42 | 0x6A42 | 27202 | 27202 | 67th reading |  | R | 1 |  |  | 0
| 0x6A43 | 0x6A43 | 27203 | 27203 | 68th reading |  | R | 1 |  |  | 0
| 0x6A44 | 0x6A44 | 27204 | 27204 | 69th reading |  | R | 1 |  |  | 0
| 0x6A45 | 0x6A45 | 27205 | 27205 | 70th reading |  | R | 1 |  |  | 0
| 0x6A46 | 0x6A46 | 27206 | 27206 | 71st reading |  | R | 1 |  |  | 0
| 0x6A47 | 0x6A47 | 27207 | 27207 | 72nd reading |  | R | 1 |  |  | 0
| 0x6A48 | 0x6A48 | 27208 | 27208 | 73rd reading |  | R | 1 |  |  | 0
| 0x6A49 | 0x6A49 | 27209 | 27209 | 74th reading |  | R | 1 |  |  | 0
| 0x6A4A | 0x6A4A | 27210 | 27210 | 75th reading |  | R | 1 |  |  | 0
| 0x6A4B | 0x6A4B | 27211 | 27211 | 76th reading |  | R | 1 |  |  | 0
| 0x6A4C | 0x6A4C | 27212 | 27212 | 77th reading |  | R | 1 |  |  | 0
| 0x6A4D | 0x6A4D | 27213 | 27213 | 78th reading |  | R | 1 |  |  | 0
| 0x6A4E | 0x6A4E | 27214 | 27214 | 79th reading |  | R | 1 |  |  | 0
| 0x6A4F | 0x6A4F | 27215 | 27215 | 80th reading |  | R | 1 |  |  | 0
| 0x6A50 | 0x6A50 | 27216 | 27216 | 81st reading |  | R | 1 |  |  | 0
| 0x6A51 | 0x6A51 | 27217 | 27217 | 82nd reading |  | R | 1 |  |  | 0
| 0x6A52 | 0x6A52 | 27218 | 27218 | 83rd reading |  | R | 1 |  |  | 0
| 0x6A53 | 0x6A53 | 27219 | 27219 | 84th reading |  | R | 1 |  |  | 0
| 0x6A54 | 0x6A54 | 27220 | 27220 | 85th reading |  | R | 1 |  |  | 0
| 0x6A55 | 0x6A55 | 27221 | 27221 | 86th reading |  | R | 1 |  |  | 0
| 0x6A56 | 0x6A56 | 27222 | 27222 | 87th reading |  | R | 1 |  |  | 0
| 0x6A57 | 0x6A57 | 27223 | 27223 | 88th reading |  | R | 1 |  |  | 0
| 0x6A58 | 0x6A58 | 27224 | 27224 | 89th reading |  | R | 1 |  |  | 0
| 0x6A59 | 0x6A59 | 27225 | 27225 | 90th reading |  | R | 1 |  |  | 0
| 0x6A5A | 0x6A5A | 27226 | 27226 | 91st reading |  | R | 1 |  |  | 0
| 0x6A5B | 0x6A5B | 27227 | 27227 | 92nd reading |  | R | 1 |  |  | 0
| 0x6A5C | 0x6A5C | 27228 | 27228 | 93rd reading |  | R | 1 |  |  | 0
| 0x6A5D | 0x6A5D | 27229 | 27229 | 94th reading |  | R | 1 |  |  | 0
| 0x6A5E | 0x6A5E | 27230 | 27230 | 95th reading |  | R | 1 |  |  | 0
| 0x6A5F | 0x6A5F | 27231 | 27231 | 96th reading |  | R | 1 |  |  | 0
| 0x6A60 | 0x6A60 | 27232 | 27232 | 97th reading |  | R | 1 |  |  | 0
| 0x6A61 | 0x6A61 | 27233 | 27233 | 98th reading |  | R | 1 |  |  | 0
| 0x6A62 | 0x6A62 | 27234 | 27234 | 99th reading |  | R | 1 |  |  | 0
| 0x6A63 | 0x6A63 | 27235 | 27235 | 100th reading |  | R | 1 |  |  | 0
| 0x6A64 | 0x6A64 | 27236 | 27236 | 101st reading |  | R | 1 |  |  | 0
| 0x6A65 | 0x6A65 | 27237 | 27237 | 102nd reading |  | R | 1 |  |  | 0
| 0x6A66 | 0x6A66 | 27238 | 27238 | 103rd reading |  | R | 1 |  |  | 0
| 0x6A67 | 0x6A67 | 27239 | 27239 | 104th reading |  | R | 1 |  |  | 0
| 0x6A68 | 0x6A68 | 27240 | 27240 | 105th reading |  | R | 1 |  |  | 0
| 0x6A69 | 0x6A69 | 27241 | 27241 | 106th reading |  | R | 1 |  |  | 0
| 0x6A6A | 0x6A6A | 27242 | 27242 | 107th reading |  | R | 1 |  |  | 0
| 0x6A6B | 0x6A6B | 27243 | 27243 | 108th reading |  | R | 1 |  |  | 0
| 0x6A6C | 0x6A6C | 27244 | 27244 | 109th reading |  | R | 1 |  |  | 0
| 0x6A6D | 0x6A6D | 27245 | 27245 | 110th reading |  | R | 1 |  |  | 0
| 0x6A6E | 0x6A6E | 27246 | 27246 | 111th reading |  | R | 1 |  |  | 0
| 0x6A6F | 0x6A6F | 27247 | 27247 | 112th reading |  | R | 1 |  |  | 0
| 0x6A70 | 0x6A70 | 27248 | 27248 | 113th reading |  | R | 1 |  |  | 0
| 0x6A71 | 0x6A71 | 27249 | 27249 | 114th reading |  | R | 1 |  |  | 0
| 0x6A72 | 0x6A72 | 27250 | 27250 | 115th reading |  | R | 1 |  |  | 0
| 0x6A73 | 0x6A73 | 27251 | 27251 | 116th reading |  | R | 1 |  |  | 0
| 0x6A74 | 0x6A74 | 27252 | 27252 | 117th reading |  | R | 1 |  |  | 0
| 0x6A75 | 0x6A75 | 27253 | 27253 | 118th reading |  | R | 1 |  |  | 0
| 0x6A76 | 0x6A76 | 27254 | 27254 | 119th reading |  | R | 1 |  |  | 0
| 0x6A77 | 0x6A77 | 27255 | 27255 | 120th reading |  | R | 1 |  |  | 0
| 0x6A78 | 0x6A78 | 27256 | 27256 | 121st reading |  | R | 1 |  |  | 0
| 0x6A79 | 0x6A79 | 27257 | 27257 | 122nd reading |  | R | 1 |  |  | 0
| 0x6A7A | 0x6A7A | 27258 | 27258 | 123rd reading |  | R | 1 |  |  | 0
| 0x6A7B | 0x6A7B | 27259 | 27259 | 124th reading |  | R | 1 |  |  | 0
| 0x6A7C | 0x6A7C | 27260 | 27260 | 125th reading |  | R | 1 |  |  | 0
| 0x6A7D | 0x6A7D | 27261 | 27261 | 126th reading |  | R | 1 |  |  | 0
| 0x6A7E | 0x6A7E | 27262 | 27262 | 127th reading |  | R | 1 |  |  | 0
| 0x6A7F | 0x6A7F | 27263 | 27263 | 128th reading |  | R | 1 |  |  | 0
| 0x6A80 | 0x6A80 | 27264 | 27264 | 129th reading |  | R | 1 |  |  | 0
| 0x6A81 | 0x6A81 | 27265 | 27265 | 130th reading |  | R | 1 |  |  | 0
| 0x6A82 | 0x6A82 | 27266 | 27266 | 131st reading |  | R | 1 |  |  | 0
| 0x6A83 | 0x6A83 | 27267 | 27267 | 132nd reading |  | R | 1 |  |  | 0
| 0x6A84 | 0x6A84 | 27268 | 27268 | 133rd reading |  | R | 1 |  |  | 0
| 0x6A85 | 0x6A85 | 27269 | 27269 | 134th reading |  | R | 1 |  |  | 0
| 0x6A86 | 0x6A86 | 27270 | 27270 | 135th reading |  | R | 1 |  |  | 0
| 0x6A87 | 0x6A87 | 27271 | 27271 | 136th reading |  | R | 1 |  |  | 0
| 0x6A88 | 0x6A88 | 27272 | 27272 | 137th reading |  | R | 1 |  |  | 0
| 0x6A89 | 0x6A89 | 27273 | 27273 | 138th reading |  | R | 1 |  |  | 0
| 0x6A8A | 0x6A8A | 27274 | 27274 | 139th reading |  | R | 1 |  |  | 0
| 0x6A8B | 0x6A8B | 27275 | 27275 | 140th reading |  | R | 1 |  |  | 0
| 0x6A8C | 0x6A8C | 27276 | 27276 | 141st reading |  | R | 1 |  |  | 0
| 0x6A8D | 0x6A8D | 27277 | 27277 | 142nd reading |  | R | 1 |  |  | 0
| 0x6A8E | 0x6A8E | 27278 | 27278 | 143rd reading |  | R | 1 |  |  | 0
| 0x6A8F | 0x6A8F | 27279 | 27279 | 144th reading |  | R | 1 |  |  | 0
| 0x6A90 | 0x6A90 | 27280 | 27280 | 145th reading |  | R | 1 |  |  | 0
| 0x6A91 | 0x6A91 | 27281 | 27281 | 146th reading |  | R | 1 |  |  | 0
| 0x6A92 | 0x6A92 | 27282 | 27282 | 147th reading |  | R | 1 |  |  | 0
| 0x6A93 | 0x6A93 | 27283 | 27283 | 148th reading |  | R | 1 |  |  | 0
| 0x6A94 | 0x6A94 | 27284 | 27284 | 149th reading |  | R | 1 |  |  | 0
| 0x6A95 | 0x6A95 | 27285 | 27285 | 150th reading |  | R | 1 |  |  | 0
| 0x6A96 | 0x6A96 | 27286 | 27286 | 151st reading |  | R | 1 |  |  | 0
| 0x6A97 | 0x6A97 | 27287 | 27287 | 152nd reading |  | R | 1 |  |  | 0
| 0x6A98 | 0x6A98 | 27288 | 27288 | 153rd reading |  | R | 1 |  |  | 0
| 0x6A99 | 0x6A99 | 27289 | 27289 | 154th reading |  | R | 1 |  |  | 0
| 0x6A9A | 0x6A9A | 27290 | 27290 | 155th reading |  | R | 1 |  |  | 0
| 0x6A9B | 0x6A9B | 27291 | 27291 | 156th reading |  | R | 1 |  |  | 0
| 0x6A9C | 0x6A9C | 27292 | 27292 | 157th reading |  | R | 1 |  |  | 0
| 0x6A9D | 0x6A9D | 27293 | 27293 | 158th reading |  | R | 1 |  |  | 0
| 0x6A9E | 0x6A9E | 27294 | 27294 | 159th reading |  | R | 1 |  |  | 0
| 0x6A9F | 0x6A9F | 27295 | 27295 | 160th reading |  | R | 1 |  |  | 0
| 0x6AA0 | 0x6AA0 | 27296 | 27296 | 161st reading |  | R | 1 |  |  | 0
| 0x6AA1 | 0x6AA1 | 27297 | 27297 | 162nd reading |  | R | 1 |  |  | 0
| 0x6AA2 | 0x6AA2 | 27298 | 27298 | 163rd reading |  | R | 1 |  |  | 0
| 0x6AA3 | 0x6AA3 | 27299 | 27299 | 164th reading |  | R | 1 |  |  | 0
| 0x6AA4 | 0x6AA4 | 27300 | 27300 | 165th reading |  | R | 1 |  |  | 0
| 0x6AA5 | 0x6AA5 | 27301 | 27301 | 166th reading |  | R | 1 |  |  | 0
| 0x6AA6 | 0x6AA6 | 27302 | 27302 | 167th reading |  | R | 1 |  |  | 0
| 0x6AA7 | 0x6AA7 | 27303 | 27303 | 168th reading |  | R | 1 |  |  | 0
| 0x6AA8 | 0x6AA8 | 27304 | 27304 | 169th reading |  | R | 1 |  |  | 0
| 0x6AA9 | 0x6AA9 | 27305 | 27305 | 170th reading |  | R | 1 |  |  | 0
| 0x6AAA | 0x6AAA | 27306 | 27306 | 171st reading |  | R | 1 |  |  | 0
| 0x6AAB | 0x6AAB | 27307 | 27307 | 172nd reading |  | R | 1 |  |  | 0
| 0x6AAC | 0x6AAC | 27308 | 27308 | 173rd reading |  | R | 1 |  |  | 0
| 0x6AAD | 0x6AAD | 27309 | 27309 | 174th reading |  | R | 1 |  |  | 0
| 0x6AAE | 0x6AAE | 27310 | 27310 | 175th reading |  | R | 1 |  |  | 0
| 0x6AAF | 0x6AAF | 27311 | 27311 | 176th reading |  | R | 1 |  |  | 0
| 0x6AB0 | 0x6AB0 | 27312 | 27312 | 177th reading |  | R | 1 |  |  | 0
| 0x6AB1 | 0x6AB1 | 27313 | 27313 | 178th reading |  | R | 1 |  |  | 0
| 0x6AB2 | 0x6AB2 | 27314 | 27314 | 179th reading |  | R | 1 |  |  | 0
| 0x6AB3 | 0x6AB3 | 27315 | 27315 | 180th reading |  | R | 1 |  |  | 0
| 0x6AB4 | 0x6AB4 | 27316 | 27316 | 181st reading |  | R | 1 |  |  | 0
| 0x6AB5 | 0x6AB5 | 27317 | 27317 | 182nd reading |  | R | 1 |  |  | 0
| 0x6AB6 | 0x6AB6 | 27318 | 27318 | 183rd reading |  | R | 1 |  |  | 0
| 0x6AB7 | 0x6AB7 | 27319 | 27319 | 184th reading |  | R | 1 |  |  | 0
| 0x6AB8 | 0x6AB8 | 27320 | 27320 | 185th reading |  | R | 1 |  |  | 0
| 0x6AB9 | 0x6AB9 | 27321 | 27321 | 186th reading |  | R | 1 |  |  | 0
| 0x6ABA | 0x6ABA | 27322 | 27322 | 187th reading |  | R | 1 |  |  | 0
| 0x6ABB | 0x6ABB | 27323 | 27323 | 188th reading |  | R | 1 |  |  | 0
| 0x6ABC | 0x6ABC | 27324 | 27324 | 189th reading |  | R | 1 |  |  | 0
| 0x6ABD | 0x6ABD | 27325 | 27325 | 190th reading |  | R | 1 |  |  | 0
| 0x6ABE | 0x6ABE | 27326 | 27326 | 191st reading |  | R | 1 |  |  | 0
| 0x6ABF | 0x6ABF | 27327 | 27327 | 192nd reading |  | R | 1 |  |  | 0
| 0x6AC0 | 0x6AC0 | 27328 | 27328 | 193rd reading |  | R | 1 |  |  | 0
| 0x6AC1 | 0x6AC1 | 27329 | 27329 | 194th reading |  | R | 1 |  |  | 0
| 0x6AC2 | 0x6AC2 | 27330 | 27330 | 195th reading |  | R | 1 |  |  | 0
| 0x6AC3 | 0x6AC3 | 27331 | 27331 | 196th reading |  | R | 1 |  |  | 0
| 0x6AC4 | 0x6AC4 | 27332 | 27332 | 197th reading |  | R | 1 |  |  | 0
| 0x6AC5 | 0x6AC5 | 27333 | 27333 | 198th reading |  | R | 1 |  |  | 0
| 0x6AC6 | 0x6AC6 | 27334 | 27334 | 199th reading |  | R | 1 |  |  | 0
| 0x6AC7 | 0x6AC7 | 27335 | 27335 | 200th reading |  | R | 1 |  |  | 0
Custom reading settings
| 0x6B00 | 0x6B00 | 27392 | 27392 | The number of bytes has been set | uint16 | R/W | 1 |  |  | 100
| 0x6B01 | 0x6B01 | 27393 | 27393 | 1st register id | uint16 | R/W | 1 |  |  | 0x3200
| 0x6B02 | 0x6B02 | 27394 | 27394 | 2nd register id | uint16 | R/W | 1 |  |  | 0x3201
| 0x6B03 | 0x6B03 | 27395 | 27395 | 3rd register id | uint16 | R/W | 1 |  |  | 0x3202
| 0x6B04 | 0x6B04 | 27396 | 27396 | 4th register id | uint16 | R/W | 1 |  |  | 0x3203
| 0x6B05 | 0x6B05 | 27397 | 27397 | 5th register id | uint16 | R/W | 1 |  |  | 0x3204
| 0x6B06 | 0x6B06 | 27398 | 27398 | 6th register id | uint16 | R/W | 1 |  |  | 0x3205
| 0x6B07 | 0x6B07 | 27399 | 27399 | 7th register id | uint16 | R/W | 1 |  |  | 0x3206
| 0x6B08 | 0x6B08 | 27400 | 27400 | 8th register id | uint16 | R/W | 1 |  |  | 0x3207
| 0x6B09 | 0x6B09 | 27401 | 27401 | 9th register id | uint16 | R/W | 1 |  |  | 0x3208
| 0x6B0A | 0x6B0A | 27402 | 27402 | 10th register id | uint16 | R/W | 1 |  |  | 0x3209
| 0x6B0B | 0x6B0B | 27403 | 27403 | 11th register id | uint16 | R/W | 1 |  |  | 0x320A
| 0x6B0C | 0x6B0C | 27404 | 27404 | 12th register id | uint16 | R/W | 1 |  |  | 0x4000
| 0x6B0D | 0x6B0D | 27405 | 27405 | 13th register id | uint16 | R/W | 1 |  |  | 0x4004
| 0x6B0E | 0x6B0E | 27406 | 27406 | 14th register id | uint16 | R/W | 1 |  |  | 0x4008
| 0x6B0F | 0x6B0F | 27407 | 27407 | 15th register id | uint16 | R/W | 1 |  |  | 0x400C
| 0x6B10 | 0x6B10 | 27408 | 27408 | 16th register id | uint16 | R/W | 1 |  |  | 0x4010
| 0x6B11 | 0x6B11 | 27409 | 27409 | 17th register id | uint16 | R/W | 1 |  |  | 0x4014
| 0x6B12 | 0x6B12 | 27410 | 27410 | 18th register id | uint16 | R/W | 1 |  |  | 0x4018
| 0x6B13 | 0x6B13 | 27411 | 27411 | 19th register id | uint16 | R/W | 1 |  |  | 0x401C
| 0x6B14 | 0x6B14 | 27412 | 27412 | 20th register id | uint16 | R/W | 1 |  |  | 0x3000
| 0x6B15 | 0x6B15 | 27413 | 27413 | 21st register id | uint16 | R/W | 1 |  |  | 0x3002
| 0x6B16 | 0x6B16 | 27414 | 27414 | 22nd register id | uint16 | R/W | 1 |  |  | 0x3004
| 0x6B17 | 0x6B17 | 27415 | 27415 | 23rd register id | uint16 | R/W | 1 |  |  | 0x3006
| 0x6B18 | 0x6B18 | 27416 | 27416 | 24th register id | uint16 | R/W | 1 |  |  | 0x3008
| 0x6B19 | 0x6B19 | 27417 | 27417 | 25th register id | uint16 | R/W | 1 |  |  | 0x300A
| 0x6B1A | 0x6B1A | 27418 | 27418 | 26th register id | uint16 | R/W | 1 |  |  | 0x300C
| 0x6B1B | 0x6B1B | 27419 | 27419 | 27th register id | uint16 | R/W | 1 |  |  | 0x300E
| 0x6B1C | 0x6B1C | 27420 | 27420 | 28th register id | uint16 | R/W | 1 |  |  | 0x3010
| 0x6B1D | 0x6B1D | 27421 | 27421 | 29th register id | uint16 | R/W | 1 |  |  | 0x3012
| 0x6B1E | 0x6B1E | 27422 | 27422 | 30th register id | uint16 | R/W | 1 |  |  | 0x3014
31th-50th are reserved. | 0x6B1F | 0x6B1F | 27423 | 27423 | 31st register id | uint16 | R/W | 1
| 0x6B20 | 0x6B20 | 27424 | 27424 | 32nd register id | uint16 | R/W | 1
| 0x6B21 | 0x6B21 | 27425 | 27425 | 33rd register id | uint16 | R/W | 1
| 0x6B22 | 0x6B22 | 27426 | 27426 | 34th register id | uint16 | R/W | 1
| 0x6B23 | 0x6B23 | 27427 | 27427 | 35th register id | uint16 | R/W | 1
| 0x6B24 | 0x6B24 | 27428 | 27428 | 36th register id | uint16 | R/W | 1
| 0x6B25 | 0x6B25 | 27429 | 27429 | 37th register id | uint16 | R/W | 1
| 0x6B26 | 0x6B26 | 27430 | 27430 | 38th register id | uint16 | R/W | 1
| 0x6B27 | 0x6B27 | 27431 | 27431 | 39th register id | uint16 | R/W | 1
| 0x6B28 | 0x6B28 | 27432 | 27432 | 40th register id | uint16 | R/W | 1
| 0x6B29 | 0x6B29 | 27433 | 27433 | 41st register id | uint16 | R/W | 1
| 0x6B2A | 0x6B2A | 27434 | 27434 | 42nd register id | uint16 | R/W | 1
| 0x6B2B | 0x6B2B | 27435 | 27435 | 43rd register id | uint16 | R/W | 1
| 0x6B2C | 0x6B2C | 27436 | 27436 | 44th register id | uint16 | R/W | 1
| 0x6B2D | 0x6B2D | 27437 | 27437 | 45th register id | uint16 | R/W | 1
| 0x6B2E | 0x6B2E | 27438 | 27438 | 46th register id | uint16 | R/W | 1
| 0x6B2F | 0x6B2F | 27439 | 27439 | 47th register id | uint16 | R/W | 1
| 0x6B30 | 0x6B30 | 27440 | 27440 | 48th register id | uint16 | R/W | 1
| 0x6B31 | 0x6B31 | 27441 | 27441 | 49th register id | uint16 | R/W | 1
| 0x6B32 | 0x6B32 | 27442 | 27442 | 50th register id | uint16 | R/W | 1
废弃
废弃
废弃
废弃
废弃
Note on F37: [线程批注] 你的Excel版本可读取此线程批注; 但如果在更新版本的Excel中打开文件，则对批注所作的任何改动都将被删除。了解详细信息: https://go.microsoft.com/fwlink/?linkid=870924 注释: The CRC32 only include the “Enable cable loss compensation” and "Cable resistance"

## Readings
Block | Start(Hex) | End(Hex) | Start(Dec) | End(Dec) | Descrption | Data type | RW | Reg | RANGE | Default | Factory value
Real-time Parameters (float) | 0x3000 | 0x3001 | 12288 | 12289 | V(Measured or Compensated) | float32 | R | 2
| 0x3002 | 0x3003 | 12290 | 12291 | Current | float32 | R | 2
| 0x3004 | 0x3005 | 12292 | 12293 | Power | float32 | R | 2
| 0x3006 | 0x3007 | 12294 | 12295 | Voltage Ripple Factor | float32 | R | 2
| 0x3008 | 0x3009 | 12296 | 12297 | Current Ripple Factor | float32 | R | 2
| 0x300A | 0x300B | 12298 | 12299 | Demand Current import | float32 | R | 2
| 0x300C | 0x300D | 12300 | 12301 | Demand Current export | float32 | R | 2
| 0x300E | 0x300F | 12302 | 12303 | Demand Power import | float32 | R | 2
| 0x3010 | 0x3011 | 12304 | 12305 | Demand Power export | float32 | R | 2
| 0x3012 | 0x3013 | 12306 | 12307 | V(Measured) | float32 | R | 2
| 0x3014 | 0x3015 | 12308 | 12309 | V(Compensated) | float32 | R | 2
| 0x3016 | 0x3017 | 12310 | 12311 | Temperature | float32 | R | 2
| 0x3018 | 0x3021 | 12312 | 12321 | Reserved | float32 | R | 10
Real-time Parameters (float) -High Speed | 0x3100 | 0x3101 | 12544 | 12545 | V(Measured or Compensated) | float32 | R | 2
| 0x3102 | 0x3103 | 12546 | 12547 | Current | float32 | R | 2
| 0x3104 | 0x3105 | 12548 | 12549 | Power | float32 | R | 2
| 0x3106 | 0x3107 | 12550 | 12551 | V(Measured) | float32 | R | 2
| 0x3108 | 0x3109 | 12552 | 12553 | V(Compensated) | float32 | R | 2
Real-time Parameters (Int) | 0x3200 | 0x3200 | 12800 | 12800 | V(Measured or Compensated) | int16 | R | 1 | 10x amplification
| 0x3201 | 0x3201 | 12801 | 12801 | Current | int16 | R | 1 | 10x amplification
| 0x3202 | 0x3202 | 12802 | 12802 | Power | int16 | R | 1 | 10x amplification
| 0x3203 | 0x3203 | 12803 | 12803 | Voltage Ripple Factor | int16 | R | 1 | 1000x amplification
| 0x3204 | 0x3204 | 12804 | 12804 | Current Ripple Factor | int16 | R | 1 | 1000x amplification
| 0x3205 | 0x3205 | 12805 | 12805 | Demand Current import | int16 | R | 1 | 100x amplification
| 0x3206 | 0x3206 | 12806 | 12806 | Demand Current export | int16 | R | 1 | 100x amplification
| 0x3207 | 0x3207 | 12807 | 12807 | Demand Power import | int16 | R | 1 | 100x amplification
| 0x3208 | 0x3208 | 12808 | 12808 | Demand Power export | int16 | R | 1 | 100x amplification
| 0x3209 | 0x3209 | 12809 | 12809 | V(Measured) | int16 | R | 1 | 10x amplification
| 0x320A | 0x320A | 12810 | 12810 | V(Compensated) | int16 | R | 1 | 10x amplification
| 0x320B | 0x320B | 12811 | 12811 | Temperature | int16 | R | 1 | 10x amplification
| 0x320C | 0x3210 | 12812 | 12816 | Reserved | int16 | R | 5
Real-time Parameters (Int) - High Speed | 0x3300 | 0x3300 | 13056 | 13056 | V(Measured or Compensated) | int16 | R | 1 | 10x amplification
| 0x3301 | 0x3301 | 13057 | 13057 | Current | int16 | R | 1 | 100x amplification
| 0x3302 | 0x3302 | 13058 | 13058 | Power | int16 | R | 1 | 100x amplification
| 0x3303 | 0x3303 | 13059 | 13059 | V(Measured) | int16 | R | 1 | 10x amplification
| 0x3304 | 0x3304 | 13060 | 13060 | V(Compensated) | int16 | R | 1 | 10x amplification
Energy & Charge | 0x4000 | 0x4003 | 16384 | 16387 | Import Energy | double | R/W | 4
| 0x4004 | 0x4007 | 16388 | 16391 | Export Energy | double | R/W | 4
| 0x4008 | 0x400B | 16392 | 16395 | Net Energy | double | R/W | 4
| 0x400C | 0x400F | 16396 | 16399 | Total Energy | double | R/W | 4
| 0x4010 | 0x4013 | 16400 | 16403 | Import Charge | double | R/W | 4
| 0x4014 | 0x4017 | 16404 | 16407 | Export Charge | double | R/W | 4
| 0x4018 | 0x401B | 16408 | 16411 | Net Charge | double | R/W | 4
| 0x401C | 0x401F | 16412 | 16415 | Total Charge | double | R/W | 4
| 0x4020 | 0x4021 | 16416 | 16417 | Energy Data Integrity Check CRC32 | uint16_t | RO | 2 | * Not available for Up Computer
Note on F53: [线程批注] 你的Excel版本可读取此线程批注; 但如果在更新版本的Excel中打开文件，则对批注所作的任何改动都将被删除。了解详细信息: https://go.microsoft.com/fwlink/?linkid=870924 注释: The CRC32 only include “Import Energy”,"Export Energy","Net Energy","Total Energy"
废弃
废弃

## System Infomation
Block | Start(Hex) | End(Hex) | Start(Dec) | End(Dec) | Descrption | Data type | Reg | RW | RANGE | Default | Factory value
Firmware Version Info | 0xF000 | 0xF001 | 61440 | 61441 | Firmware version | uint8_t | 2 | R | In the format like x.yz (e.g., 1.01) | determined by firmware | determined by firmware
| 0xF002 | 0xF004 | 61442 | 61444 | Firmware release date | uint8_t | 3 | R | In the format like YYMMDD (221228) | determined by firmware | determined by firmware
| 0xF005 | 0xF005 | 61445 | 61445 | Firmware patch number | uint8_t | 1 | R | In the format like xy (01) | determined by firmware | determined by firmware
Reboot | 0xF000 | 0xF001 | 61440 | 61441 | Reset for update | uint8_t | 2 | W | In the format 0xA1 0x58 0x58 0xA1 |  |  | write with 0x6A
Boot Loader | 0xF010 | 0xF011 | 61456 | 61457 | bootloader product string | uint8_t | 2 | R | Main Controller：MDB1 (temply, Not Fixed） Remote Display: MDB2 (temply，Not Fixed） | determined by boot | determined by boot
| 0xF012 | 0xF013 | 61458 | 61459 | bootloader version | uint8_t | 2 | R | In the format like x.yz (e.g., 1.01) | determined by boot | determined by boot
| 0xF014 | 0xF016 | 61460 | 61462 | bootloader release date | uint8_t | 3 | R | ASCII, in format of YYMMDD (e.g. 241108) | determined by boot | determined by boot
| 0xF017 | 0xF017 | 61463 | 61463 | Bootloader Patch Number | uint8_t | 1 | R | In the format like xy (01) | determined by boot | determined by boot
Product Info | 0xF040 | 0xF04F | 61504 | 61519 | Serial Number | uint8_t | 16 | RW | In the format like DE+YY(+30)+MM+abcd(0001-9999)(e.g., DE55050001)
| 0xF050 | 0xF051 | 61520 | 61521 | Hardware version | uint8_t | 2 | RW | In the format like x.yz (e.g., 1.01) | 1.01 | 1.01
| 0xF052 | 0xF052 | 61522 | 61522 | Function Model Type | uint16_t | 1 | RW | 0: default (Not permit for DC321) 1: AcuDC-321 | 1 | 1 | UI display: AcuDC-301
| 0xF053 | 0xF053 | 61523 | 61523 | Voltage Input Type | uint16_t | 1 | RW | 0 for 1500V | 0 | 0 | UI display: 1000V
| 0xF054 | 0xF054 | 61524 | 61524 | Current Input Type | uint16_t | 1 | RW | 0 for A1 (1500A) | 0 | 0 | UI display: 650A
| 0xF055 | 0xF055 | 61525 | 61525 | Power Supply Type | uint16_t | 1 | RW | 1 for P2 | 1 | 1 | UI display: DC
| 0xF056 | 0xF065 | 61526 | 61541 | MAC address | uint8_t | 16 | RW
| 0xF066 | 0xF066 | 61542 | 61542 | Device Type | uint16_t | 1 | RW
| 0xF067 | 0xF067 | 61543 | 61543 | Remote Display Mode | uint16_t | 1 | RW | 0: Master Mode; 1: Slave Mode：Upgrade 2:Slave Mode：FCT test | 0
| 0xF068 | 0xF06F | 61544 | 61551 | Reserved | uint16_t | 8 | RW
| 0xF070 | 0xF070 | 61552 | 61552 | Hardware Patch Number | uint8_t | 1 | RW | In the format like xy (01) | 00
| 0xF071 | 0xF07A | 61553 | 61562 | Reserved 2 | uint16_t | 10 | RW
| 0xF07B | 0xF09B | 61563 | 61595 | public key | uint8 | 33 | RW | raw public key, totally 65 bytes
| 0xF09C | 0xF0AB | 61596 | 61611 | private key | uint8 | 16 | W | raw private key,totally 32 bytes
| 0xF0AC | 0xF0AD | 61612 | 61613 | CRC32 for product info | uint32 | 2 | RO | * Not available for Up Computer
※ write with function code 0x6A in factory
Meter System Status | 0xF0B4 | 0xF0B4 | 61620 | 61620 | Integrity Check Result | uint16 | 1 | RO | BIT0 (1- There is Boot firmware Integrity check failure. 0- There isn't) BIT1 (1- There is App firmware integrity check failure. 0- There isn't) BIT2 (1- There is Product integrity check failure. 0- There isn't) BIT3 (1- There is ADC calibration integrity check failure. 0- There isn't) BIT4 (1- There is Measure configuration integrity check failure.0- There isn't) BIT5 (1- There is Energy data integrity check failure. 0- There isn't) BIT6 (1- There is Echilog data integrity check failure. 0- There isn't) BIT7 (1- There is Transaction data integrity check failure. 0- There isn't) Other BITS - Reserved
| 0xF0B5 | 0xF0B5 | 61621 | 61621 | System Mode | uint16 | 1 | RO | 0x0001 - Production Mode 0x0002 - Assembly Mode 0x0003 - Normal Mode Others - Reserved
| 0xF0B6 | 0xF0B6 | 61622 | 61622 | Fatal Error Status | uint16 | 1 | RO | BIT0 (1- Data integrity failure happened, 0- No data integrity failure.) BIT1 (1- Log storge is exhausted, 0- Log storge is not exhausted.) BIT2 (1 - RTC works not correctly, 0- RTC works correctly) BIT3 (1 - ADC works not correctly, 0- ADC works correctly) BIT4 (1 - Storage failure occurred, 0- No storage failure.) Other bits - Reserved.
| 0xF0B7 | 0xF0B7 | 61623 | 61623 | Assembly mode enter command | uint16 | 1 | WR | Write 0xAA55,enter the assembly mode. Don't response any others write value. Self clear.
Public Key
| 0xF0C0 | 0xF0ED | 61632 | 61677 | public key | uint8 | 46 | RW | public key of .der format, totally 91 bytes |  |  | use 0x6A to write; 0x3 to read
R
R
R
废弃

## Max, Min
Block | Start(Hex) | End(Hex) | Start(Dec) | End(Dec) | Descrption | Data type | RW | Reg | RANGE | Default | Factory value
Max | 0x4136 | 0x4137 | 16694 | 16695 | V Max Value | Float32 | R | 2
| 0x4138 | 0x413E | 16696 | 16702 | V Max Timestamp | INT | R | 7 |  |  |  | 7th reg is the millesecond value
| 0x413F | 0x4140 | 16703 | 16704 | I Max Value | Float32 | R | 2
| 0x4141 | 0x4147 | 16705 | 16711 | I Max Timestamp | INT | R | 7
| 0x4148 | 0x4149 | 16712 | 16713 | P Max Value | Float32 | R | 2
| 0x414A | 0x4150 | 16714 | 16720 | P Max Timestamp | INT | R | 7
| 0x4151 | 0x4152 | 16721 | 16722 | Ripple factor U Max Value | Float32 | R | 2
| 0x4153 | 0x4159 | 16723 | 16729 | Ripple factor U Max Timestamp | INT | R | 7
| 0x415A | 0x415B | 16730 | 16731 | Ripple factor I Max Value | Float32 | R | 2
| 0x415C | 0x4162 | 16732 | 16738 | Ripple factor I Max Timestamp | INT | R | 7
Demand Max | 0x4163 | 0x4164 | 16739 | 16740 | Demand I import Max Value | Float32 | R | 2
| 0x4165 | 0x416B | 16741 | 16747 | Demand I import Max Timestamp | INT | R | 7
| 0x416C | 0x416D | 16748 | 16749 | Demand I export Max Value | Float32 | R | 2
| 0x416E | 0x4174 | 16750 | 16756 | Demand I export Max Timestamp | INT | R | 7
| 0x4175 | 0x4176 | 16757 | 16758 | Demand P import Max Value | Float32 | R | 2
| 0x4177 | 0x417D | 16759 | 16765 | Demand P import Max Timestamp | INT | R | 7
| 0x417E | 0x417F | 16766 | 16767 | Demand P export Max Value | Float32 | R | 2
| 0x4180 | 0x4186 | 16768 | 16774 | Demand P export Max Timestamp | INT | R | 7
Min | 0x4187 | 0x4188 | 16775 | 16776 | V Min Value | Float32 | R | 2
| 0x4189 | 0x418F | 16777 | 16783 | V Min Timestamp | INT | R | 7
| 0x4190 | 0x4191 | 16784 | 16785 | I Min Value | Float32 | R | 2
| 0x4192 | 0x4198 | 16786 | 16792 | I Min Timestamp | INT | R | 7
| 0x4199 | 0x419A | 16793 | 16794 | P Min Value | Float32 | R | 2
| 0x419B | 0x41A1 | 16795 | 16801 | P Min Timestamp | INT | R | 7
| 0x41A2 | 0x41A3 | 16802 | 16803 | Ripple factor U Min Value | Float32 | R | 2
| 0x41A4 | 0x41AA | 16804 | 16810 | Ripple factor U Min Timestamp | INT | R | 7
| 0x41AB | 0x41AC | 16811 | 16812 | Ripple factor I Min Value | Float32 | R | 2
| 0x41AD | 0x41B3 | 16813 | 16819 | Ripple factor I Min Timestamp | INT | R | 7
废弃

## Datalog（废弃）
Block | Start(Hex) | End(Hex) | Start(Dec) | End(Dec) | Descrption | Data type | RW | Reg | RANGE | Default | Factory value
Datalog 1, 2, 3 setting | 0x1100 | 0x1100 | 4352 | 4352 | Log1 #Registers # Sectors |  | R/W | 1
| 0x1101 | 0x1101 | 4353 | 4353 | Log1 Interval |  | R/W | 1
| 0x1102 | 0x1133 | 4354 | 4403 | Log1 register #1-#50 identifier |  | R/W | 50
| 0x1134 | 0x1134 | 4404 | 4404 | Datalog mode selection |  | R/W | 1
| 0x1135 | 0x1135 | 4405 | 4405 | Start month, year |  | R/W | 1
| 0x1136 | 0x1136 | 4406 | 4406 | Start day, hour |  | R/W | 1
| 0x1137 | 0x1137 | 4407 | 4407 | Start minute, second |  | R/W | 1
| 0x1138 | 0x1138 | 4408 | 4408 | End month, year |  | R/W | 1
| 0x1139 | 0x1139 | 4409 | 4409 | End day, hour |  | R/W | 1
| 0x113A | 0x113A | 4410 | 4410 | End minute, second |  | R/W | 1
| 0x113B | 0x113B | 4411 | 4411 | Clear Data Log 1 |  | R/W | 1
| 0x113C | 0x1177 | 4412 | 4471 | Log2 setting |  | R/W | 60
| 0x1178 | 0x11B3 | 4472 | 4531 | Log3 setting |  | R/W | 60
Datalog 1 Sector | 0x1400 | 0x1400 | 5120 | 5120 | first sector, current sector | word | R | 1
| 0x1401 | 0x1401 | 5121 | 5121 | current address | word | R | 1
| 0x1402 | 0x1402 | 5122 | 5122 | full (low byte) | word | R | 1 | 0: not full 1: sector is full
Datalog 2 Sector | 0x1410 | 0x1410 | 5136 | 5136 | first sector, current sector | word | R | 1
| 0x1411 | 0x1411 | 5137 | 5137 | current address | word | R | 1
| 0x1412 | 0x1412 | 5138 | 5138 | full (low byte) | word | R | 1 | 0: not full 1: sector is full
Datalog 3 Sector | 0x1420 | 0x1420 | 5152 | 5152 | first sector, current sector | word | R | 1
| 0x1421 | 0x1421 | 5153 | 5153 | current address | word | R | 1
| 0x1422 | 0x1422 | 5154 | 5154 | full (low byte) | word | R | 1 | 0: not full 1: sector is full
Datalog 4 Sector | 0x1430 | 0x1430 | 5168 | 5168 | first sector, current sector | word | R | 1
| 0x1431 | 0x1431 | 5169 | 5169 | current address | word | R | 1
| 0x1432 | 0x1432 | 5170 | 5170 | full (low byte) | word | R | 1 | 0: not full 1: sector is full
Datalog 4 setting | 0x1500 | 0x1500 | 5376 | 5376 | Log4 #Registers # Sectors |  | R/W | 1 | 1 -50；0-64 | 0 | 0
| 0x1501 | 0x1501 | 5377 | 5377 | Log4 Interval |  | R/W | 1 | 1 - 1440 | 0 | 0
| 0x1502 | 0x1533 | 5378 | 5427 | Log4 register #1-#50 identifier |  | R/W | 50 | Parameter selection address range: 0x3000, 0x3002, 0x3004, 0x3006, 0x3008, 0x300A, 0x300C, 0x300E, 0x3010, 0x3012, 0x3014 | 0 | 0
| 0x1534 | 0x1534 | 5428 | 5428 | Datalog mode selection |  | R/W | 1 | 0 - 2
| 0x1535 | 0x1535 | 5429 | 5429 | Start month, year |  | R/W | 1 | Low byte: month, high byte: year (keep consistence with Data Log123) | 0 | 0
| 0x1536 | 0x1536 | 5430 | 5430 | Start day, hour |  | R/W | 1 | Low byte: hour, high byte: day
| 0x1537 | 0x1537 | 5431 | 5431 | Start minute, second |  | R/W | 1 | Low byte: second, high byte: minute
| 0x1538 | 0x1538 | 5432 | 5432 | End month, year |  | R/W | 1 | Low byte: month, high byte: year
| 0x1539 | 0x1539 | 5433 | 5433 | End day, hour |  | R/W | 1 | Low byte: hour, high byte: day
| 0x153A | 0x153A | 5434 | 5434 | End minute, second |  | R/W | 1 | Low byte: second, high byte: minute
| 0x153B | 0x153B | 5435 | 5435 | Clear Data Log 4 |  | R/W | 1 | write 1 to clear
Datalog 1 reading | 0x6000 | 0x6000 | 24576 | 24576 | Record type, reserved |  | R/W | 1
| 0x6001 | 0x6001 | 24577 | 24577 | Record count of each window , window state |  | R/W | 1
| 0x6002 | 0x6003 | 24578 | 24579 | OFFSET |  | R/W | 2
| 0x6004 | 0x60CC | 24580 | 24780 | Window |  | R | 201
Datalog 1 status | 0x6100 | 0x6101 | 24832 | 24833 | Max Records | uint32 | R | 2
| 0x6102 | 0x6103 | 24834 | 24835 | Used Records | uint32 | R | 2
| 0x6104 | 0x6104 | 24836 | 24836 | Recordsize | uint16 | R | 1
| 0x6105 | 0x6105 | 24837 | 24837 | Log Availability | uint16 | R | 1
| 0x6106 | 0x6108 | 24838 | 24840 | First Record Timestamp | word | R | 3
| 0x6109 | 0x610B | 24841 | 24843 | Last Record Timestamp | word | R | 3
| 0x610C | 0x610D | 24844 | 24845 | Record Index | uint32 | R | 2
Datalog 2 status | 0x6200 | 0x6201 | 25088 | 25089 | Max Records | uint32 | R | 2
| 0x6202 | 0x6203 | 25090 | 25091 | Used Records | uint32 | R | 2
| 0x6204 | 0x6204 | 25092 | 25092 | Recordsize | uint16 | R | 1
| 0x6205 | 0x6205 | 25093 | 25093 | Log Availability | uint16 | R | 1
| 0x6206 | 0x6208 | 25094 | 25096 | FirstRecordTimestamp | word | R | 3
| 0x6209 | 0x620B | 25097 | 25099 | LastRecordTimestamp, | word | R | 3
| 0x620C | 0x620D | 25100 | 25101 | Record Index | uint32 | R | 2
Datalog 3 status | 0x6300 | 0x6301 | 25344 | 25345 | Max Records | uint32 | R | 2
| 0x6302 | 0x6303 | 25346 | 25347 | Used Records | uint32 | R | 2
| 0x6304 | 0x6304 | 25348 | 25348 | Recordsize | uint16 | R | 1
| 0x6305 | 0x6305 | 25349 | 25349 | Log Availability | uint16 | R | 1
| 0x6306 | 0x6308 | 25350 | 25352 | FirstRecordTimestamp | word | R | 3
| 0x6309 | 0x630B | 25353 | 25355 | LastRecordTimestamp, | word | R | 3
| 0x630C | 0x630D | 25356 | 25357 | Record Index | uint32 | R | 2
Datalog 4 reading | 0x6400 | 0x6400 | 25600 | 25600 | Record type, reserved |  | R/W | 1 | 0 - 3 （3 for Data Log4） | 0 | 0
| 0x6401 | 0x6401 | 25601 | 25601 | Record count of each window +window state |  | R/W | 1 | High byte: Record number of window (setting) Low byte: window state (0x0B: data effective,0xFF: data not effective, 0xAA: Data Log on clearing, 0xBB: Data Log clearing end)
| 0x6402 | 0x6403 | 25602 | 25603 | OFFSET |  | R/W | 2
| 0x6404 | 0x64CC | 25604 | 25804 | Window |  | R | 201 | Data Log format: record number (4bytes)+ time stamps (6bytes) + [data1~dataN] (6Nbytes) + CRC (2bytes) Data format: parameter1 average value, parameter2 average value... parameter1 max value, parameter2 max value, parameter1 min value, parameter2 min value...
Datalog 4 status | 0x6500 | 0x6501 | 25856 | 25857 | Max Records | uint32 | R | 2 | 3904 - 232960
| 0x6502 | 0x6503 | 25858 | 25859 | Used Records | uint32 | R | 2 | 1 - 232960
| 0x6504 | 0x6504 | 25860 | 25860 | Recordsize | uint16 | R | 1 | 18 - 1074
| 0x6505 | 0x6505 | 25861 | 25861 | Log Availability（reserve） | word | R | 1
| 0x6506 | 0x6508 | 25862 | 25864 | First Record Time stamp | word | R | 3
| 0x6509 | 0x650B | 25865 | 25867 | Last Record Time stamp | word | R | 3
| 0x650C | 0x650D | 25868 | 25869 | Record Index | uint32 | R | 2
note: max records of Data Log4 = 65536 / (number of register * 6 + 12) * sector number (fixed to 64 at present)
整个 Sheet 废弃

## TransactionLog
Block | Start(Hex) | End(Hex) | Start(Dec) | End(Dec) | Description | Data type | RW | Reg | RANGE | Default | Factory value
Settings | 0x5000 | 0x5000 | 20480 | 20480 | Identification Status | uint16 | RW | 1 | 0 = False 1 = True
| 0x5001 | 0x5001 | 20481 | 20481 | Identification Level | uint16 | RW | 1 | 0 = “NONE”1 = “HEARSAY”2 = “TRUSTED” 3 = “VERIFIED” 4 = “CERTIFIED”5 = “SECURE”6 = “MISMATCH”7 = “INVALID” 8 = “OUTDATED”9= “UNKNOWN”
| 0x5002 | 0x5002 | 20482 | 20482 | Identification Flags 1 | uint16 | RW | 1 | 0 = RFID_NONE (Default) 1 = RFID_PLAIN 2 = RFID_RELATED 3 = RFID_PSK
| 0x5003 | 0x5003 | 20483 | 20483 | Identification Flags 2 | uint16 | RW | 1 | 0 = OCPP_NONE (Default) 1 = OCPP_RS 2 = OCPP_AUTH 3 = OCPP_RS_TLS 4 = OCPP_AUTH_TLS 5 = OCPP_CACHE 6 = OCPP_WHITELIST 7 = OCPP_CERTIFIED
| 0x5004 | 0x5004 | 20484 | 20484 | Identification Flags 3 | uint16 | RW | 1 | 0 = ISO15118_NONE (Default) 1 = ISO15118_PNC
| 0x5005 | 0x5005 | 20485 | 20485 | Identification Flags 4 | uint16 | RW | 1 | 0 = PLMN_NONE (Default) 1 = PLMN_RING 2 = PLMN_SMS
| 0x5006 | 0x5006 | 20486 | 20486 | Identification Type | uint16 | RW | 1 | 0 = NONE (Default) 1 = DENIED 2 = UNDEFINED 3 = ISO14443 4 = ISO15693 5 = EMAID 6 = EVCCID 7 = EVCOID 8 = ISO7812 9 = CARD_TXN_NR 10 = CENTRAL 11 = CENTRAL_1 12 = CENTRAL_2 13 = LOCAL 14 = LOCAL_1 15 = LOCAL_2 16 = PHONE_NUMBER 17 = KEY_CODE
| 0x5007 | 0x501A | 20487 | 20506 | Identification Data | uint8 | RW | 20 | max 40 chars
| 0x501B | 0x502E | 20507 | 20526 | Tariff Text | uint8 | RW | 20 | max 40 chars
| 0x502F | 0x502F | 20527 | 20527 | transaction timeout in seconds | uint16 | RW | 1 | default 0 (disabled)， range: 3600~64800
| 0x5030 | 0x5030 | 20528 | 20528 | CT: Charge Point Identification Type | uint16 | RW | 1 | 0: EVSEID 1: CBIDC
| 0x5031 | 0x5044 | 20529 | 20548 | CI: Charge Point Identification | uint8 | RW | 20 | 40 chars: Default = all byte set with 0x00
| 0x5045 | 0x5045 | 20549 | 20549 | Time zone shift | int16 | RW | 1 | Max = 1440 Min = -1440, default = 0
| 0x5046 | 0x5049 | 20550 | 20553 | UTC timestamp for synchronization | uint64 | RW | 4 | Default = 0
Block And Id | 0x5100 | 0x5100 | 20736 | 20736 | current Block id | word | R | 1
| 0x5101 | 0x5101 | 20737 | 20737 | current address | word | R | 1
| 0x5102 | 0x5103 | 20738 | 20739 | min Transaction Id | uint32 | R | 2
| 0x5104 | 0x5105 | 20740 | 20741 | max Transaction Id | uint32 | R | 2
| 0x5106 | 0x5106 | 20742 | 20742 | full (low byte) | word | R | 1 | 0: not full 1: sector is full
Reading | 0x5200 | 0x5200 | 20992 | 20992 | OCMF Command | uint16 | RW | 1 | 0: do nothing "B" = start transaction "E" = End of transaction "A" = Abort transaction
| 0x5201 | 0x5201 | 20993 | 20993 | OCMF file generate state | uint16 | R | 1 | 0 = not ready 1 = ready to generate 2 = generating 3 = ready to read
| 0x5202 | 0x5205 | 20994 | 20997 | transaction Time | uint64 | R | 4
| 0x5206 | 0x5206 | 20998 | 20998 | timeSync status | uint16 | R | 1 | 0 = Time not synchronized 1 = Time synchronized
| 0x5207 | 0x520A | 20999 | 21002 | Transaction Import Energy in kWh | double | R | 4
| 0x520B | 0x520E | 21003 | 21006 | Transaction Export Energy in kWh | double | R | 4
| 0x520F | 0x5210 | 21007 | 21008 | Duration (in second) | uint32 | R | 2
| 0x5211 | 0x5211 | 21009 | 21009 | transaction running status | uint16 | R | 1 | 0 = idle 1 = charging 2 = out of power 3 = end normally 4 = canceled 5 = end with overtime
| 0x5212 | 0x5213 | 21010 | 21011 | OFFSET | uint32 | RW | 2
| 0x5214 | 0x5215 | 21012 | 21013 | ocmf last transaction id | uint32 | R | 2
| 0x5216 | 0x5217 | 21014 | 21015 | ocmf first transaction id | uint32 | R | 2
| 0x5218 | 0x5219 | 21016 | 21017 | transactionLog id | uint32 | RW | 2 | transactionLog id for reading
| 0x521A | 0x521A | 21018 | 21018 | transactionLog OCMF file size | uint16 | R | 1
| 0x521B | 0x5472 | 21019 | 21618 | transactionLog OCMF data | uint8 | R | 600
| 0x5473 | 0x5474 | 21619 | 21620 | transaction id in charging | uint32 | R | 2
| 0x5475 | 0x5478 | 21621 | 21624 | Start time for charging | uint64 | R | 4
Status | 0x5500 | 0x5501 | 21760 | 21761 | Max Records | uint32 | R | 2
| 0x5502 | 0x5503 | 21762 | 21763 | Used Records | uint32 | R | 2
| 0x5504 | 0x5504 | 21764 | 21764 | Recordsize | uint16 | R | 1 | fixed as 62

## Echilog
Block | Start(Hex) | End(Hex) | Start(Dec) | End(Dec) | Description | Data type | RW | Reg | RANGE | Default | Factory value
Echilog Status | 0x5700 | 0x5700 | 22272 | 22272 | Max Records | uint16 | RO | 1 | Fixed as 6000
| 0x5701 | 0x5701 | 22273 | 22273 | Record Index | uint16 | RO | 1
| 0x5702 | 0x5702 | 22274 | 22274 | Echilog full status | uint16 | RO | 1 | 0x0000 - Not full 0x0001 - Full Others - Reserved
| 0x5703 | 0x5703 | 22275 | 22275 | Record Number | uint16 | RO | 1
Echilog Reading | 0x5800 | 0x5800 | 22528 | 22528 | The Echilog ID to to read | uint16_t | RW | 1 | 1~6000
| 0x5801 | 0x5801 | 22529 | 22529 | The Echilog number need to read | uint16_t | RW | 1 | Fixed as 0x01
| 0x5802 | 0x5802 | 22530 | 22530 | The Echilog read window status | uint16_t | RO | 1 | 0x0B - The data windown is ready 0xFF - The request data is invalid;
| 0x5803 | 0x5803 | 22531 | 22531 | The signatured bytes size of the Echilog to be read | uint16_t | RO | 1
| 0x5804 | 0x587B | 22532 | 22651 | Echilog data window | uint8_t | RO | 120
Firmware CRC
| 0x5910 | 0x5911 | 22800 | 22801 |  | uint32_t | RO | 2 | Boot Firmware Stored CRC32
| 0x5849 | 0x584A | 22601 | 22602 |  | uint32_t | RO | 2 | Boot Firmware Calculated CRC32
| 0x584A | 0x584B | 22602 | 22603 |  | uint32_t | RO | 2 | App Firmware Stored CRC32
| 0x584B | 0x584C | 22603 | 22604 |  | uint32_t | RO | 2 | App Firmware Calculated CRC32

## Calibration
Block | Start(Hex) | End(Hex) | Start(Dec) | End(Dec) | Descrption | Data type | RW | Reg | RANGE | Default | Factory value
Calibration Commands | 0xF204 | 0xF204 | 61956 | 61956 | enter calibration mode (0x6A write) | word | W | 1 |  |  |  | with parameter 0x00, 0xaa
| 0xF205 | 0xF205 | 61957 | 61957 | start calibration (0x6A write) | word | W | 1 |  |  |  | with parameter 0x00, 0xaa
| 0xF206 | 0xF206 | 61958 | 61958 | Reserved1 | word | W | 1 |  |  |  | with parameter 0x00, 0xaa
| 0xF207 | 0xF207 | 61959 | 61959 | Reserved2 | word | W | 1 |  |  |  | with parameter 0x00, 0xaa
| 0xF208 | 0xF208 | 61960 | 61960 | return the calibration result | word | R | 1
| 0xF209 | 0xF209 | 61961 | 61961 | exit calibration mode (0x6A write） | word | W | 1 |  |  |  | with parameter 0x00, 0x01
Calibration Values | 0xF210 | 0xF211 | 61968 | 61969 | V offset and I offset MSB | uint32 | RO | 2
| 0xF212 | 0xF213 | 61970 | 61971 | V offset and I offset LSB | uint32 | RO | 2
| 0xF214 | 0xF215 | 61972 | 61973 | V gain and I gain MSB | uint32 | RO | 2
| 0xF216 | 0xF217 | 61974 | 61975 | V gain and I gain LSB | uint32 | RO | 2
| 0xF218 | 0xF219 | 61976 | 61977 | Calibration Integrity Check CRC32 | uint16_t | RO | 2 | * Not available for Up Computer
Note on F15: [线程批注] 你的Excel版本可读取此线程批注; 但如果在更新版本的Excel中打开文件，则对批注所作的任何改动都将被删除。了解详细信息: https://go.microsoft.com/fwlink/?linkid=870924 注释: The CRC32 include all the calibration data